Introduction

The decontamination of ultrasound probes is essential to reduce the risk of contracting a healthcare-acquired infection.

This toolkit has been assembled in consultation with clinical experts with backgrounds in infection prevention and instrument reprocessing.

The objective in developing this toolkit has been to provide a resource regarding infection prevention during the use and reprocessing of ultrasound probes.

This toolkit will support you to locate the ultrasound probes within your organisation, as well as guide you on how to measure the risks associated with the procedure for which the probe has been used.

Table of Contents

Tool 1: Part A - Locate	3

Tool 1: Part B - Profile	3

Tool 2: Algorithm	3

Tool 3: Example Risk Assessment	4

Tool 4: Policy Development Framework	4

Foreword	4

Approach and Principles	5

Funding	5

Important Note	5

Disclaimer	5

Acronyms	6

Tool 1: Part A – Locate	8

Tool 1: Part B – Profile	10

Tool 1: Part C – Action Plan	14

References	16

# Tool 1: Part A - Locate

This list of strategies can help locate ultrasound machines in your facility. Locating ultrasound machines can be challenging, as they may move around, be uncatalogued, or be unknown to personnel responsible for probe use and reprocessing.

# Tool 1: Part B - Profile

After locating ultrasound machines with the Locate tool, this audit tool containing ready to use forms and observation checklists can be used to observe and assess procedure-specific ultrasound SOP and practice. First, list of ultrasound systems present in the deparment, record the probes used and related prcedures, then assess the documentation in place before auditing the actual practice.

# Tool 2: Algorithm

This tool is organised by department and provides a range of typical procedures that may be encountered in that department. Probe use and reprocessing requirements based on national guidelines (NHS Scotland, HSE Ireland, SCoR/BMUS) and international standards (ESR, ECMUS and WFUMB) are presented as a decision making algorithm. The flow chart can be printed out and displayed throughout office and procedure rooms and used as a quick reference chart for healthcare workers to determine whether a practice is compliant with available guidelines. An editable chart allows further adjustment following approved local processes.

Tool 3: Example Risk Assessment

This tool contains four editable templates designed to guide the assessment of potential hazards that may be encountered during the use and reprocessing (cleaning, disinfection, storage) of ultrasound probes. A sample risk matrix is provided with further instructions for completion.  A facility should aim to mitigate all significant harm to the lowest risk rating. If that is not possible, the existing workflow and/or products should be reconsidered.

# Tool 4: Policy Development Framework

This tool is a policy development framework designed to help develop infection prevention policies for all settings where ultrasound technology is used and probes are reprocessed. It can be used to develop a universal hospital policy or a department-specific policy and has been developed based on major guidelines, standards and evidence-based scientific literature.

# Foreword

This UK and Ireland toolkit has been adapted from the original U.S. toolkit, which was assembled in consultation with North American clinical experts with backgrounds in infection prevention and instrument reprocessing. The objective in developing this adapted toolkit has been to provide a relevant resource regarding infection prevention during the use and reprocessing of ultrasound probes in the United Kingdom and Ireland.

UK-specific adaptions have been made by:

Linda Cooper Decontamination Project Lead (Northern Ireland)

Niamh O’Callaghan Hospital Sterile Services Clinical Nurse Manager & Decontamination Lead (Ireland)

John Prendergast Authorised Engineer (Decontamination)(Wales)

Helen Campbell Decontamination Consultant (England)

Anthony Sullivan External Decontamination Services Manager (England)

Elizabeth Collins Clinical Lead Infection Prevention (England)

Claire Jones-Manning Deputy Head of Operations & Decontamination Lead (England)

The following individuals contributed to the development of the original US toolkit:

Sue Barnes RN CIC FAPIC

Roy Boukidjian MSN PHN CIC

Ruth Carrico PhD DNP APRN FSHEA CIC

Robert Garcia BS MT CIC

Sylvia Garcia-Houchins RN MBA CIC

Amy Nichols RN MBA CIC FAPIC

Rose Seavey MBA BS RN CNOR CRCST CSPDT

John Whelan BSN RN

# Approach and Principles

These tools have been assembled based on best practice national and regional guidelines 1-10 with the goal of reducing infection risks associated with ultrasound use in the interest of patient safety. They are intended to guide the development of institutional policies and procedures, with an understanding that each institution can vary in the patient groups cared for, settings, types of care provided.

# Funding

The development of this toolkit has been supported by Nanosonics Ltd. Nanosonics Ltd strives to improve the safety of patients, clinics, their staff and the environment by transforming the way infection prevention practices are understood and conducted and introducing innovative technologies that deliver improved standards of care.

# Important Note

The information, materials and opinions contained in this toolkit have been prepared for general information purposes only, are not intended to constitute clinical or other professional advice and should not be relied on or treated as a substitute for such advice on the subject matter.

Please consider this guidance in light of the reader’s own professional training and judgement and specific regulations, guidelines, policies and procedures of each region, institution and department. This toolkit does not replace directions contained in a manufacturer instructions for use (IFUs) documents, nor does it replace institutional policies/workflows, but is intended to be used in conjunction with them. This Important Note applies to all parts of this toolkit.

# Disclaimer

Nanosonics manufactures and supplies automated high-level disinfection systems for ultrasound probes. Whilst all commercially reasonable efforts have been taken to prepare this toolkit, Nanosonics will not accept any liability or responsibility of any kind for reliance on the information, materials and opinions contained in this toolkit, including for any death or injury to persons.  Nanosonics makes no warranties, representations or undertakings about the content (including, without limitation, the quality, accuracy, completeness or fitness for any particular purpose of such content). To the fullest extent permitted by law, Nanosonics will not be liable for any loss or damage (including any consequential, incidental, special or indirect damages) whatsoever or for any loss of data, loss arising from interruption to business, loss of profits or other damage or losses of any third party arising out of or in any way relating to errors, omissions, or inaccurate or out-of-date information in this document. Any and all use or reliance on this toolkit is at your own risk.

Healthcare professionals should undertake their own enquiries as necessary and exercise professional judgement when assessing the needs of the particular patient. Copies of the journal articles and references referred to in the toolkit can be provided upon request.

# Acronyms

AEd: Authorised Engineer (Decontamination)

BMUS: The British Medical Ultrasound Society

ECMUS: European Committee for Medical Ultrasound Safety

ESR: European Society of Radiology

HCO: Health Care Organisation

HLD: High-Level Disinfection

HSE Ireland: Health Service Executive Ireland

HTM: Health Technical Memorandum

IFU: Instructions of Use

ILD: Intermediate-Level Disinfection

IPC: Infection Prevention and Control

IQ: Installation Qualification

LLD: Low-Level Disinfection

MHRA: Medicines & Healthcare products Regulatory Agency

NHS Scotland: National Health Services Scotland

OQ: Operation Qualification

PPE: Personal Protective Equipment

PPM: Planned Preventive Maintenance

P.Q.: Performance Qualification

RIMDS: Reusable Invasive Medical Devices

SCoR: Society and College of Radiographers

SOP: Standard Operating Procedure

SSD: Sterile Services Department

WFUMB: European Federation of Societies for Ultrasound in Medicine and Biology

# Tool 1: Part A – Locate

Objective: Locate the departments/clinics using ultrasound at your facility.

This list of strategies can help locate ultrasound machines in your facility. Locating ultrasound machines can be a challenging task, as they may move around, be uncatalogued, or may be unknown to personnel with responsibility for probe use and reprocessing.

The options listed can be used to determine where ultrasound is used in order to facilitate follow up on any issues related to infection prevention. The different options below may be relevant based on the type and size of your facility. More than one strategy may be used in attempting to locate ultrasound machines. Strategies are ranked in order of assumed effectiveness.

Department name: double click to type here

# Tool 1: Part B – Profile

Objective: Profile policy and practice for each ultrasound procedure

at your facility.

After locating ultrasound machines with the Locate tool, this audit tool can be used to observe and assess procedure-specific ultrasound policy and practice. The page can be printed, and ultrasound procedures for each facility area can be used for individual observation sessions. First, record the procedure and department, then work through the profile form. The questions guide the user through their policy, and an observation checklist is provided to record actual practice. The tool leads the user to an action plan (part C) if there is no policy or if there are discrepancies between policy, practice and guidelines.

Tool 1: Part B – Profile

Profile policy and practice for each ultrasound procedure at your facility.

a) Department and ultrasound systems

| Department | Date | Assessor / Auditor | Operator(s) present |
| --- | --- | --- | --- |

|  | Ultrasound machines identification | Manufacturer(s) |
| --- | --- | --- |
| 1 |  |  |
| 2 |  |  |
| 3 |  |  |
| 4 |  |  |
| 5 |  |  |
| 6 |  |  |
| 7 |  |  |
| 8 |  |  |
| 9 |  |  |
| 10 |  |  |

b) Probes and procedures

| Probe name | Probe type (surface / endo-cavitary) | Probe ID* | Probe Serial Number | Probe Manufacturer | Procedure(s) | Spaulding Classification (non critical / semi-critical / critical) |
| --- | --- | --- | --- | --- | --- | --- |

* ID number given by the HCO

c) Documentation

| Questions | Answers |  |  | Comments |
| --- | --- | --- | --- | --- |
| Does the facility/department have SOP(s) for reprocessing ultrasound probes? | □ Yes | □ No | □ Not specified |  |
| Do the operators have access to the US probe(s) IFU? | □ Yes | □ No | □ Not specified |  |
| Is the reprocessing system manual? | □ Yes | □ No | □ Not specified |  |
| Is the reprocessing system automated? | □ Yes | □ No | □ Not specified |  |
| If yes to the automated question above: Has the automated reprocessor been validated, serviced and tested in line with national and international standards and following the manufacturer's recommendation? | □ Yes | □ No | □ Not specified |  |
| Where are the validation and service reports stored? | □ Yes | □ No | □ Not specified |  |

d) Probe Reprocessing Procedure Audit:

Procedure name/type: _____________________________________________

Spaulding classification: 	□ Non-critical       □ Semi-critical         □ Critical

| SOP name | SOP number | SOP revision date |
| --- | --- | --- |

| Questions | Answers |  |  | Comments |
| --- | --- | --- | --- | --- |
| Is the SOP compliant with the relevant guidelines standards and manufacturer’s IFU (compatibility)? | □ Yes | □ No | □ Not specified |  |

Observe Practice: How is the probe reprocessed and used for this procedure?

| Step | Observations |  |  |  | Comments |
| --- | --- | --- | --- | --- | --- |
| Post-procedure cleaning (bed-side) | □ No | □ Yes | Method used: | □ Not specified |  |
| Containment of contaminated probe | □ No | □ Yes | Method used: | □ Not specified |  |
| Cleaning | □ No | □ Yes | Method used: | □ Not specified |  |
| Disinfection type | □ LLD | □ HLD | Method used: | □ Not specified |  |
| Disinfection method | □ Manual | □ Automated | Method used: | □ Not specified |  |
| Terminal process step | □ LLD | □ HLD | □ sterilisation | □ Not specified |  |
| Cover use | □ None | □ Single-use non-sterile | □ Sterile | □ Not specified |  |
| Gel use | □ Multi-use bottle | □ Single-use non-sterile | □ Single-use sterile | □ Not specified |  |
| Multi-use Gel max usage | □ 1 day | □ > 1 day | Please specify: | □ Not specified |  |
| Reprocessing Traceability | □ No traceability | □ Manual | □ Electronic | □ Not specified |  |
| Post processing storage | □ No | □ Yes | Method used: | □ Not specified |  |
| Time limit defined? | □ No | □ Yes | Time limit: | □ Not specified |  |
| Other steps(s) (please specify) |  |  |  |  |  |

Assess Practice

| Questions | Answers |  |  | Comments |
| --- | --- | --- | --- | --- |
| Is the observed practice fully compliant with the SOP? | □ Yes | □ No | □ Not specified |  |
| Where are the training records stored? | Please specify: |  | □ Not specified |  |
| Are the operators trained for this procedure? | □ Yes | □ No | □ Not specified |  |
| Type of training | □ Internal | □ External | □ Not specified |  |

Audit Conclusion

Comments

If your SOP and practice are compliant: File this form, ensure ongoing training for all users.

If your SOP needs updating and/or users are not trained according to your SOP: Go to Part C.

# Tool 1: Part C – Action Plan

Objective: Address any deviations identified in Part B – Profile.

Create an itemised action plan. The action plan should have a review date to assess its effectiveness and re-audit if necessary.

To ensure compliance, continuous improvement and effectiveness, the action plan should be reviewed at least annually.

Audit location: ___________________________________________________________

| Action Plan |  |  |  | Progress Update | Re-audit |  |
| --- | --- | --- | --- | --- | --- | --- |
| Auditor/assessor: Date: |  |  |  | Auditor/assessor: Date: | Auditor/assessor: Date: |  |
| Deviations identified | Actions required | Person responsible | Priority High Medium Low | Status | Outcome | Action effectiveness (How do you prove that the actions have been effective?) |

Action Plan Example

| Action Plan |  |  |  | Progress Update | Re-audit |  |
| --- | --- | --- | --- | --- | --- | --- |
| Auditor/assessor: Date: |  |  |  | Auditor/assessor: Date: | Auditor/assessor: Date: |  |
| Deviations identified | Actions required | Person responsible | Priority High Medium Low | Status | Outcome | Action effectiveness (How do you prove that the actions have been effective?) |
| SOP to be created | 1.Review IFU | XX | High | 1.IFU reviewed | SOP conform with IFU |  |
|  | 2.Create a standardised procedure | XX | High | 2. SOP created | SOP is available in folder for all staff | Perform process audits once a year |
|  | 3.Communicate SOP to all relevant staff | Y.Y. | Medium | 3. Awareness training rolled out | All relevant staff have been asked to describe the SOP |  |
|  | 4. Update training folders | Y.Y. | Medium | Staff training folders to be updated |  |  |

# References

- Society and College of Radiographers and British Medical Ultrasound Society (2017). "Guidelines For Professional Ultrasound Practice." 127.

- Health Service Executive (HSE) Quality Improvement Division (2017). HSE Guidance for Decontamination of Semi‐critical Ultrasound Probes; Semi‐invasive and Non‐invasive Ultrasound Probes. Document: QPSD-GL-028-1.

- Health Protection Scotland (HPS) (2016). NHS Scotland Guidance for Decontamination of Semi-Critical Ultrasound Probes; Semi-invasive and Non-invasive Ultrasound Probes.

- Nyhsen, C. M., et al. (2017). "Infection prevention and control in ultrasound - best practice recommendations from the European Society of Radiology Ultrasound Working Group." Insights Imaging 8(6): 523-535.

- Kollmann, C. and K. Salvesen (2017). "Best Practice recommendations for cleaning and disinfection of ultrasound transducers whilst maintaining transducer integrity." European Committee for Medical Ultrasound Safety (ECMUS).

- Abramowicz, J. S., et al. (2017). "Guidelines for Cleaning Transvaginal Ultrasound Transducers Between Patients." Ultrasound Med Biol 43(5): 1076-1079.

- Department of Health (U.K.) (2016). Health Technical Memorandum 01-06 Part C Operational management. Department of Health (U.K.), Department of Health (U.K.),. C: 72.

- Department of Health (U.K.) (2016). Health Technical Memorandum 01-06 Part D Validation and verification (including storage/drying cabinets). Department of Health (U.K.): 49.

- Welsh Health Technical Memorandum (WHTM) (2014). WHTM 01-06 - Decontamination of flexible endoscopes Part C: Operational management, NHS Wales Shared Services Partnership – Specialist Estates Services: 74.

- Rutala, W. A., et al. (2008). Guideline for Disinfection and Sterilisation in Healthcare Facilities (Updated 15 February 2017).

- Spaulding EH (1968). Chemical disinfection of medical and surgical materials. Disinfection, sterilisation, and preservation. Lawrence C, Block SS. Philadelphia (P.A.), Lea & Febiger: 517-531.

- Masood J, Voulgaris S, Awogu O, Younis C, Ball AJ, Carr TW. Condom perforation during transrectal ultrasound-guided (TRUS) prostate biopsies: a potential infection risk. International urology and nephrology. 2007;39(4):1121-4.

- Amis S, Ruddy M, Kibbler CC, Economides DL, MacLean AB. Assessment of condoms as probe covers for transvaginal sonography. J Clin Ultrasound. 2000;28(6):295-8.

- Milki AA, Fisch JD. Vaginal ultrasound probe cover leakage: implications for patient care. Fertil Steril. 1998;69(3):409-11.

- Storment JM, Monga M, Blanco JD. Ineffectiveness of latex condoms in preventing contamination of the transvaginal ultrasound transducer head. South Med J. 1997;90(2):206-8.

- Rooks VJ, Yancey MK, Elg SA, Brueske L. Comparison of probe sheaths for endovaginal sonography. Obstet Gynecol. 1996;87(1):27-9.

- Hignett M, Claman P. High rates of perforation are found in endovaginal ultrasound probe covers before and after oocyte retrieval for in vitro fertilisation-embryo transfer. J Assist Reprod Genet. 1995;12(9):606-9.

- Fritz S, Hust MH, Ochs C, Gratwohl I, Staiger M, Braun B. Use of a latex cover sheath for transesophageal echocardiography (TEE) instead of regular disinfection of the echoscope? Clinical Cardiology. 1993;16(10):737-40.

- Abdelfattah R, Aljumaah S, Alqahtani A, Althawadi S, Barron I, Almofada S. Outbreak of Burkholderia cepacia bacteraemia in a tertiary care centre due to contaminated ultrasound probe gel. Journal of Hospital Infection. 2018;98(3):289-94.

- Shaban RZ, Maloney S, Gerrard J, Collignon P, Macbeth D, Cruickshank M, et al. Outbreak of health care-associated Burkholderia cenocepacia bacteremia and infection attributed to contaminated sterile gel used for central line insertion under ultrasound guidance and other procedures. American journal of infection control. 2017;45(9):954-58.

- Cheng A, Sheng WH, Huang YC, Sun HY, Tsai YT, Chen ML, et al. Prolonged postprocedural outbreak of Mycobacterium massiliense infections associated with ultrasound transmission gel. Clinical Microbiology and Infection. 2016;22(4):382.e1-.e11.

- Ofstead CL, Wetzler HP, Snyder AK, Horton RA. Endoscope reprocessing methods: a prospective study on the impact of human factors and automation. Gastroenterology nursing: the official journal of the Society of Gastroenterology Nurses and Associates. 2010;33(4):304-11.

- HSE Ireland – Guidance on the Implementation of the Health and Safety Competent Authority – Exemption 03/2016 on the Carriage of Uncleaned Reusable Invasive Medical Devices by Road (QPSD-GL-021-1).

- Public Health England. Good infection prevention practice: using ultrasound gel. 2021;GW-1865.

- Health Technical Memorandum (HTM) 01-01 on the management and decontamination of surgical instruments (medical device) used in acute care. Part A: Management and provision – July 2016.

- HSE Standards and Recommended Practices for Central Decontamination Units QPSD-D-003-2.1 V2.1. Part 3 Recommended Practices, 13.4 Procedure Pg 92.

- MHRA Medical Device Alert – Reusable transoesophageal echocardiography, transvaginal and transrectal ultrasound probes (transducers). MDA/2012/037 28 June 2021.

- REGULATION (E.U.) 2017/745 OF THE EUROPEAN PARLIAMENT AND OF THE COUNCIL of 5 April 2017 on medical devices, amending Directive 2001/83/E.C., Regulation (E.C.) No 178/2002 and Regulation (C.E.) No 1223/2009 and repealing Council Directives 09/385/EEC and 93/42/EEC.

- The Carriage of Dangerous Goods (Amendment) Regulations 2019.